## pipette_calibration
Worksheet for exercise on measurement uncertainty in pipetting | [6 empty] | Fill the cellls highlighted YELLOW
| Date | 2013-02-07 |  |  | Yes/No
| Balance checked and found OK |  |  |  | yes
| Pipette checked for cleanliness and absence of chips |  |  |  | yes
| Pipette filler function checked and found OK |  |  |  | yes
| Results: | Weight of water +tare | Tare | weight of nominal 25 mL water by pipette | calculated volume
| 1 | 42.1906 | 17.323 | 24.8676 | 24.9422 |  | Table of corrections
| 2 | 42.1649 | 17.3257 | 24.8392 | 24.9137 |  | 17 | 1.0023
| 3 | 42.1512 | 17.2913 | 24.8599 | 24.9345 |  | 18 | 1.0025
| 4 | 42.1835 | 17.4026 | 24.7809 | 24.8552 |  | 19 | 1.0026
| 5 | 42.1831 | 17.2978 | 24.8853 | 24.96 |  | 20 | 1.0028
| 6 | 42.1826 | 17.3037 | 24.8789 | 24.9535 |  | 21 | 1.003
| 7 | 42.1673 | 17.2722 | 24.8951 | 24.9698 |  | 22 | 1.0032
| 8 | 42.1862 | 17.2883 | 24.8979 | 24.9726 |  | 23 | 1.0034
| 9 | 42.1717 | 17.2985 | 24.8732 | 24.9478 |  | 24 | 1.0036
| 10 | 42.1796 | 17.3006 | 24.879 | 24.9536 |  | 25 | 1.0038
|  |  |  | Mean | 24.9403 |  | 26 | 1.0041
|  |  |  | Standard deviation | 0.03446 |  | 27 | 1.0043
|  |  |  | Relative standard deviation | 0.0013817
|  |  | Start, C | End | Mean
|  | Water temperature | 22.4 | 21 | 21.7
|  | Water density | 0.99701 | Correction | 1.003 | 1.0036

## burette_calibration
Worksheet for exercise on burette calibration
| Date | 2012-01-01 |  |  | Yes/No
| Balance checked and found OK |  |  |  | no
| Burette checked for cleanliness and absence of chips |  |  |  | no
|  | Weight of water +tare | Tare | weight of water | calculated volume | Deviation
First Run | 10 | 1 | 0 | 1 | 1.0038 | 8.9962
| 20 | 2 | 0 | 2 | 2.0076 | 17.9924
| 30 | 3 | 0 | 3 | 3.0114 | 26.9886
| 40 | 4 | 0 | 4 | 4.0152 | 35.9848
| 50 | 5 | 0 | 5 | 5.019 | 44.981 | Table of corrections
|  |  |  | 0 |  |  | 17 | 1.0023
Second run | 10 | 1 | 0 | 1 | 1.0038 | 8.9962 | 18 | 1.0025
| 20 | 2 | 0 | 2 | 2.0076 | 17.9924 | 19 | 1.0026
| 30 | 3 | 0 | 3 | 3.0114 | 26.9886 | 20 | 1.0028
| 40 | 4 | 0 | 4 | 4.0152 | 35.9848 | 21 | 1.003
| 50 | 5 | 0 | 5 | 5.019 | 44.981 | 22 | 1.0032
|  |  | Average: |  |  |  | 23 | 1.0034
|  |  |  | 10 | 1.0038 | 8.9962 | 24 | 1.0036
|  |  |  | 20 | 2.0076 | 17.9924 | 25 | 1.0038
|  |  |  | 30 | 3.0114 | 26.9886 | 26 | 1.0041
|  |  |  | 40 | 4.0152 | 35.9848 | 27 | 1.0043
|  |  |  | 50 | 5.019 | 44.981
|  |  | Start, C | End | Mean
|  | Water temperature | 25 | 25 | 25
|  | Water density | 0.99621 |  | 1.0038 | 1.0038